## Slide 1: X-Ray & Gamma Radiation: Measurement Best Practices Max Fomitchev-Zamilov, Ph.D. Maximus Energy Corporation founder@maximus.energy www.maximus.energy

## Slide 2: Philosophy of Discovery
Form a hypothesis, describe confirmation and falsification criteria;
Subdue your ego: the truth is more important;
Question everything including your own conclusions;
Look for errors and you will find them;
Invite critique from everybody, especially from peers.

## Slide 3: Philosophy of Measurement
Understand how your measurement devices work;
Achieve repeatability of measurements;
Study and eliminate systematic errors;
Employ alternative measurement techniques;
Use bank controls: the effect must disappear;
Use broad controls: the effect must increase/decrease when you vary critical parameters;
Do not forget the statistical analysis.

## Slide 4: Gamma Radiation Detectors
Spectroscopy-Grade NaI(Tl) Scintillators

## Slide 5: Gamma Spectrometer Schematics
NaI(Tl) Crystal
PMT
Voltage Divider
HV Power Supply
Computer
MCA
Amplifier
Preamp
Cable

## Slide 6: Do Not Use “Black Box” Systems
“Black box” systems are not suitable for scientific studies
Interference
Incorrect Measurements

## Slide 7: Gamma Spectrometer Simplified
NaI(Tl) Crystal
PMT
Voltage Divider
HV Power Supply
ADC
Computer

## Slide 8: GAMMA-PRO NaI(Tl) Spectrometer
http://maximus.energy/index.php/product/gamma-pro/

## Slide 9: PulseCounter Software
http://maximus.energy/index.php/software/

## Slide 10: NaI Detector Properties
Sensitivity: depends on size (width and thickness) of scintillator crystal;
Energy Range: from 10 keV to tens of MeV (depending on scintillator crystal thickness);
Resolution: energy-dependent; good crystals have ~30% @ 30 keV, ~7% @ 662 keV, 5% @ 1.3 MeV.
Central hole / well for compact sample measurement with 4π coverage.

## Slide 11: NaI Detector Efficiency
Exponentially decreases with gamma energy
Exponentially increases with scintillator crystal thickness
Has a maximum around ~100 keV
Absolute efficiency is 2-7% on average (depending on gamma spectrum)

## Slide 12: NaI Detector Resolution
Peak width decreases with the increase in gamma energy from ~30% @ 50 keV to ~5% @ > 1 MeV
NaI detectors are not very accurate for x-ray (<100 keV) spectroscopy due to poor resolution

## Slide 13: Detectors: Area of Applicability
SDD – high resolution (~2% @ 6 keV), low sensitivity, max range from 0.2 to 100 keV; optimal range from 1 to 30 keV.
CdTe/CdTeZn/CZT – good resolution (~4% @ 662 кэВ), low sensitivity, max range from 1 keV to 1.5 MeV; optimal range from 1.5 to 100 keV.
NaI(Tl) – average resolution (~7-8% @ 662 keV), high sensitivity, max range from 10 keV 10 MeV; optimal range from 30 keV to 3 MeV.
HPGe – ultra-high resolution (0.2% @ 1.3 MeV), low sensitivity, useful range from 3 keV to 10 MeV (depending on type).

## Slide 14: SDD
High Resolution: ~2% @ 6 keV
Low Sensitivity
Optimal Range: 1 – 30 keV
Maximum Range: 0.2 – 100 keV
Peak efficiency @ ~25 keV

## Slide 15: CdTe / CdZnTe / CZT
Good Resolution: ~4% @ 662 keV
Low Sensitivity
Optimal Range: 1.5 – 100 keV
Maximum Range: 1 keV – 1.5 MeV
Peak Efficiency @ ~10-60 keV

## Slide 16: NaI, CsI
Average Resolution: ~7-8% @ 662 keV
High Sensitivity
Optimal Range: 30 keV – 3 MeV
Maximum Range: 10 keV – 10 MeV
Peak Efficiency @ ~100-200 keV

## Slide 17: Germanium Detectors / HPGe
Ultra-High Resolution: ~0.2% @ 1.3 MeV
Low Sensitivity
Optimal Range Depends on Detector Type
Common Range: 3 keV – 3 MeV
Peak Efficiency @ ~200 keV

## Slide 18: Typical NaI(Tl) Detector Signal

## Slide 19: Typical NaI(Tl) Detector Signal
Base Level: 0V
Trigger Level (Discriminator): -10 mV
Rise Time : ~2 μs
Pulse Width: ~20 μs
Pulse Amplitude: -200 mV

## Slide 20
Exaggerated Signal
Trigger Level (Discriminator)
Noise
Useful Signal

## Slide 21: Natural NaI Gamma Background

## Slide 22: Calibration Sources
137Cs, 60Co and other calibration sources can be purchased at @ www.imagesco.com;
One has to have at least two calibration peaks in the energy range of interest;
One should check calibration before every experiment (no “stale” calibration).

## Slide 23: Sources of Systematic Errors
Intrinsic noise produces false peaks in the low spectrum channels;
Periodic EM interference may produce false peaks across all spectrum channels, may shift/distort energies of spectral peaks;
Magnetic field interferes with PMT operation;
Concurrent gammas cannot be resolved and register as a single gamma of cumulative energy;
Temperature change results in shift of spectral peaks;
Unintended shielding with human body or equipment may drastically affect counts and spectral peaks;
Strong gamma sources produce x-ray fluorescence (XRF) peaks;
Calibration non-linearity limits accuracy of energy resolution;
Contamination (including with radon) may produce natural radioactivity peaks (e.g. 214Pb);
Natural background may vary significantly, especially when only a few samples are taken.

## Slide 24: Controlling Systematic Errors
Intrinsic noise is mitigated by increasing the trigger level, which results in loss of low-energy gammas;
Periodic EM interference is mitigated by shielding and ground loop elimination;
Magnetic field is mitigated by screening the PMT with mu-metal and situating detectors away from magnets and high current sources;
Concurrent gammas can be resolved by switching from slow NaI to fast EJ-232 scintillator;
Temperature change can be mitigated by thermostat or by frequent recalibration;
Shielding can be mitigated by fixing the detector position and preventing human or equipment movement during measurement;
Calibration non-linearity can be mitigated by employing multiple calibration sources;
Contamination with atmospheric radon can be mitigated by operating in a hermetically sealed environment;
Natural background should be checked periodically and sampled sufficiently (at least 20 samples).

## Slide 25: Random Processes
We are dealing with random processes. That is why we cannot draw conclusions after making only 1-2 (or a handful) of measurements;
Too few measurements can yield odd patterns, including some pretty incredible ones;
“Non-repeatable” results are (most likely) just odd samples drawn from random distribution.

## Slide 26: Statistical Analysis
Sample the background: ≥20 samples
Sample the experiment: ≥20 samples
Calculate P-value: if the value is less than 5% it is customary to consider the difference between the experiment and the background to be “statistically significant”.
If the P-value changes significantly with the addition or removal of samples – you do not have enough samples in your data set.

## Slide 27: How to Calculate P-Value in Excel

## Slide 28: Statistical Analysis in PulseCounter
Линия свинца

## Slide 29: X-Ray Liminescence
Lead Peak

## Slide 30: Shielding
I stood next to the detector and
screened some background

## Slide 31: Amplification of the Effect
Increasing the number of measurements should result in P-value decrease (i.e. improved statistics);
Using more sensitive (in the energy range of interest) detector should result in the increased signal;
Using lead screen for background shielding should increase the signal;
Using metal filters to reduce or eliminate gamma rays of certain energies should affect the signal accordingly.

## Slide 32: “Gold Standard” Publication Criterion
A publishable paper must have:
Detailed description of measurement equipment and measurement methodology;
Description of calibration process;
Multiple / periodic background measurements;
Repeatable results (somewhat repeatable is OK, 100% repeatability is not required);
Multiple measurements pertaining to the effect (must have significant sample);
Statistical analysis and P-value calculation when comparing measurements;
Analysis of systematic errors: show that you have gone at reasonable length to eliminate them;
‘Blank’ control experiment with null-result;
Control experiment that increases or decreases the effect;
Mandatory secondary confirmation e.g. using principally different measurement equipment and/or alternative measurement methodology;
Peer review, especially from experts in measurement techniques and data interpretation.

## Slide 33: Useful Resources
http://maximus.energy/index.php/product-category/gamma/
http://maximus.energy/index.php/2020/05/09/search-for-alpha-decay-acceleration/
https://www.picotech.com/products/oscilloscope
https://www.amptek.com/products/sdd-x-ray-detectors-for-xrf-eds/xr-100sdd-silicon-drift-detector