## Slide 1: Neutron Radiation: Measurement Best Practices Max Fomitchev-Zamilov, Ph.D. Maximus Energy Corporation founder@maximus.energy www.maximus.energy

## Slide 2: Neutron Detection Basics
Neutrons can be broadly classified as:
  Fast: Ek > 1 MeV,
  Slow: Ek ~ 10 eV,
  Thermal: Ek ~ 0.025 eV.
Each class requires it’s own special detection technique.
Neutrons are usually rare: statistical analysis is necessary!
Natural neutron background can vary wildly: statistical analysis in necessary!
Detection probability exponentially increases with the decrease in neutron energy.

## Slide 3: Neutron Absorption Cross-Section
A thermal neutron is ~10 000 easier to detect than a fast one!
Helium-3 is almost 10 times more sensitive to thermal neutrons than Lithium-6.

## Slide 4: Neutron Moderators
Hydrogen
  Water, oil, organics
  High-density ultrahigh molecular weight polyethylene (HDPE)
  Absorbs thermal neutrons
Deuterium
  Heavy water
  Expensive
Graphite
  Must be boron-free
Beryllium
  Expensive, dust is very toxic (hence difficult to machine)

## Slide 5: Moderator Positioning
Moderator
Compact Source
Detector
Moderator
Extended Source
Detector
Optimal moderator thickness depends on neutron energy:
  Fast neutrons: 5-10 cm of HDPE
  Thermal neutrons: moderator not needed! It will absorb them
  Cold neutrons: will be absorbed by environment, even by the air

## Slide 6: Common Neutron Detector Types
Thermal neutrons
  Proportional counters: 3He, 10BF3, 10B
  Corona mode counters: typically stronger single than proportional counters
Fast neutrons: 4He, CH4, liquid scintillators
Crystal scintillators: 6Li glass, NaI(Tl + Li)
Bubble detectors
Track detectors: CR39
Solid-state: 6LiF6, 10B over silicon diode

## Slide 7: Proportional Counters 3He / 10BF3 / 10B
Very high thermal neutron sensitivity
Sensitive to electromagnetic interference
Sensitive to gamma radiation, gamma rejection is a must
Old soviet SHM-18, SI-19N are good for research

## Slide 8: 6Li Glass Crystal Scintillators
High thermal neutron sensitivity
Not as sensitive to EM interference as proportional counters
Sensitive to gamma radiation, gamma rejection is a must

## Slide 9: Bubble Detectors
Made by Bubble Technology Industries (BTI), Canada
Made for different energy ranges (e.g. one can order a 6-vial spectrometer)
Insensitive to electromagnetic interference, x-ray or gamma radiation
Inexpensive: ~$300 ea.
Low neutron sensitivity (~1000 less than 3He)
Useful life: 1-2 years
Sensitive to vibration, shock and heat (cause bubbles)

## Slide 10: CR39 Plastic Track Detector
Widely available from Landauer, USA
“Gold standard” for publication
Very inexpensive: ~$20 ea. under dosimetry subscription service
Do not provide real-time counts: require etching and counting under microscope (automated under Landauer subscription service)
Very low sensitivity (10,000 less than 3He)

## Slide 11: Neutron Activation
Irradiate indium, gold, silver or europium foil
Acquire gamma spectrum
  Requires large (3-5”) NaI or germanium (HPGe) detector in a lead shield

## Slide 12: Proportional Counter Properties
3He gas pressure: 0.5 – 20 bar
  SI-19N/SNM-18: 2-3 bar
Dimensions: diameter 5-50 mm, length up to 2 m
Bias voltage: 800-1,500V
Pulse amplitude: 10-100 mV
Pulse rise time: 1-20 μs
Background: 1-5 CPM*
Efficiency: 60-90%*
Sensitivity: depends on geometry, gas pressure:
  CPS/nv – counts per second per neutron per 1 cm2.

## Slide 13: General Counter Diagram
Bias Power Supply
Counter
Discriminator
Preamp
Detector
Short Cable
Long Cable
Filter

## Slide 14: Do Not Use “Black Box” Systems
“Black box” systems are not suitable for scientific studies
Interference
Incorrect Measurements

## Slide 15: Simplified Detector Schematics
Bias Power Supply
Preamp
Detector
Filter
Computer

## Slide 16: NEUTRON-PRO System
http://maximus.energy/index.php/product/neutron-pro/

## Slide 17: NEUTRON-BANK 12 x 12”

## Slide 18: PulseCounter Software
http://maximus.energy/index.php/software/

## Slide 19: 3He Counter: Proportional Mode
Baseline: 0V
Trigger level: -5 mV
Pulse rise time: 12 μs
Pulse width: ~100 μs
Pulse amplitude: 25 mV

## Slide 20
3He Counter: Corona Mode
Trigger level (discriminator)
Corona noise
Purported neutron pulse

## Slide 21: Thermal Neutron Spectrum
Thermal peak: 764 keV
Triton continuum to 191 keV
Proton continuum to 573 keV

## Slide 22: Gamma Rejection
Neutron spectrum
Gamma peak
Gamma peak creeps to the right as the intensity of x-rays increases

## Slide 23: Gamma Rejection
Neutron detector overwhelmed by x-rays

## Slide 24: Neutron Check Source
Nuclespot P2042 5 mCi 210Po alpha source available at www.amstat.com
Cover with beryllium and obtain a check source on the order of few thousand of neutrons per second!
People also use 241Am, 239Pu from smoke detectors
Few-Ci 252Cf calibration source is the best but it is expensive and requires an NRC license

## Slide 25: Systematic Errors
Proportional counters are very sensitive to EM interference!
  Intrinsic noise may cause false low-energy peaks and may affect the counting rate!
  Periodic EM noise (e.g. power supply noise) may add a parasitic periodic component to the counting rate!
Gamma radiation requires careful gamma rejection.
Bias or amplifier drift may impact the counting rate!
Natural neutron background may vary wildly, especially when too few measurements are taken.
Human body is a good moderator, may affect the counting rate.

## Slide 26: Random Processes
We are dealing with random processes. That is why we cannot draw conclusions after making only 1-2 (or a handful) of measurements;
Too few measurements can yield odd patterns, including some pretty incredible ones;
“Non-repeatable” results are (most likely) just odd samples drawn from random distribution.

## Slide 27: Statistical Analysis
Sample the background: ≥20 samples
Sample the experiment: ≥20 samples
Calculate P-value: if the value is less than 5% it is customary to consider the difference between the experiment and the background to be “statistically significant”.
If the P-value changes significantly with the addition or removal of samples – you do not have enough samples in your data set.

## Slide 28: How to Calculate P-Value in Excel

## Slide 29: Statistical Analysis: Example #1
P-value = 0.0007 !!!
Interesting… But too few (<30) measurements to be convincing

## Slide 30: Statistical Analysis: Example #2
Interesting… But this is just a “lucky” pattern

## Slide 31: Statistical Analysis: Example #3
One cannot discard lower “experiment” counts as errors

## Slide 32: Neutron Detection Best Practices
Configuration
  Cover the detector in lead foil to suppress gamma
  Use moderator: 2-4” of HDPE is usually enough
  Place the detector and all the electronics in a metal enclosure (Faraday cage)
  Employ battery power to avoid ground loops
  Record the detector signal for off-line analysis and pulse shape audit
Calibration
  Use a neutron check-source (e.g. 5 mCi Po-Be) to check the detector: the detector must respond to the source
  Use a gamma source (e.g. 10 μCi 137Cs) to verify gamma rejection: the detector must not respond to the source

## Slide 33: Neutron Detection Best Practices
Experiment
  Place the detector as close to the purported neutron source as possible
  Use gamma detector for gamma intensity & spectrum monitoring
Statistical Analysis
  Check background regularly: record 20+ measurements each time
  Record 20+ “experiment” measurements each time
  Calculate P-value
  Statistics will improve with the increase in the number of measurements
Secondary Means of Confirmation
  Use bubble detector to verify / confirm the neutron flux
  If the neutron flux appears large use indium foil activation to verify / confirm the neutron flux

## Slide 34: Useful Resources
http://maximus.energy/index.php/product/neutron-pro/
http://bubbletech.ca/product_type/bubble-detector-products/
https://amstat.com/esd-ionizer/nuclespot-alpha-ionizer/